## Slide 1
Band | Limit | Tones | Bare floor |  | Carpet on it |  | Final state | 
Hz | ± dB | per band | worst mic, dB | cells in limit | worst mic, dB | cells in limit | worst mic, dB | cells in limit
250 | 2.5 | 3 | 5.71 | 52 % | 0.59 | 100 % | 1.88 | 97 %
315 | 2.5 | 8 | 2.60 | 52 % | 1.41 | 91 % | 0.89 | 83 %
400 | 2.5 | 8 | 2.05 | 70 % | 0.87 | 86 % | 1.41 | 80 %
500 | 2.5 | 8 | 2.25 | 68 % | 1.09 | 91 % | 1.99 | 91 %
630 | 2.5 | 8 | 1.64 | 82 % | 1.37 | 85 % | 1.49 | 86 %
800 | 2.0 | 8 | 1.42 | 72 % | 1.14 | 83 % | 1.12 | 83 %
1000 | 2.0 | 9 | 0.89 | 82 % | 0.89 | 98 % | 0.57 | 99 %
1250 | 2.0 | 8 | 0.72 | 90 % | 0.64 | 98 % | 1.02 | 99 %
1600 | 2.0 | 8 | 1.39 | 93 % | 1.53 | 93 % | 1.15 | 99 %
2000 | 2.0 | 8 | 0.98 | 94 % | 0.87 | 98 % | 0.77 | 99 %
2500 | 2.0 | 8 | 0.81 | 86 % | 0.80 | 98 % | 0.75 | 93 %
5000* | 2.0 | 4 | 1.77 | 95 % | 2.32 | 93 % | 1.23 | 91 %
6300* | 3.0 | 5 | 1.36 | 100 % | 1.62 | 100 % | 1.77 | 100 %
WHERE WE STAND
Against the hemi-anechoic ISO values
Limits: ±2.5 dB up to 630 Hz, ±2.0 dB from 800 to 5000 Hz, ±3.0 dB at 6300 Hz. Colour: worst microphone as a share of the limit (yellow at the limit, red at twice).
Bands over the limit (of 13)
Bare floor 2 → carpet 1 → final 0
Cells in limit (13 bands)
80 % → 93 % → 92 %
Room error below 3 kHz
2.05 → 1.32 → 1.39 dB
Bare floor: 24 Sep, wedges off the floor, flat foam base only. Carpet: same afternoon, source unmoved. Final: 25 Sep (other day, source position not documented). * Partial bands: tones above 5 kHz only; 3–5 kHz left out because of the loudspeaker. ISO values are an analogue, not a qualification. Hemi limits: Winker & Stahnke 2016, Table 1 (quoting ISO 3745:2012).
Notes: Hemi-anechoic tolerance values of ISO 3745:2012 as printed in Winker and Stahnke 2016, Table 1 (and Kayhan 2008 p. 13): +-2.5 dB up to 630 Hz, +-2.0 dB for 800 to 5000 Hz, +-3.0 dB from 6300 Hz. They are looser than the anechoic values (+-1.5 / +-1.0 / +-1.5) used elsewhere in this deck. Our floor is treated, so the anechoic row remains the stricter and primary reading; this slide shows where the room stands if judged as a hemi-anechoic one. Bare floor = 2026-09-24/floor-no-wedges (14:05, flat foam base, wedges off), carpet = 2026-09-24/chaotic-carpet (14:32), final = 2026-09-25/carpet-reordered. Bare floor and carpet are the same-source pair; the final state is another day. Same method as the colour table slide: worst microphone band-mean deviation from the arc mean, and the share of tone x microphone cells inside the limit.